Uncertainties of measurement

There are two methods for determination of uncertainty.

Method A – based on mathematical statistics

Method B – based on information given by measuring devices

Method A

This method uses mathematical statistics. An average value is calculated first by the formula.

N is number of measurements

xi – i-th sample

Then the standard uncertainty equal to the standard deviation of average value is given by

Method B

For successful calculation of this method we need some data from measuring devices, mostly their precision or resolution. If the resolution of the device is Δ (which means one basic division at a needle device, for example), then the B type uncertainty is given by

Caliper

Δ = 20 um,

Micrometer

Δ = 10 um				uB= 2,9 um

Needle measuring device

This device is characterized by a class of precision CP and its range R.

Class B uncertainty is given by uB= (2x(R*CP)/100) / sqrt(12) or uB= ((R*CP)/100) / sqrt(3)

Example: CP= 0,5 % ; R= 600 mA

Digital measuring device

Precision is typically given by expression p% of rdg + n digits

uB= ((p*rdg)/100 + ls digit value) / sqrt(3)

Example: digital multimeter, full range is 19,99 V; p= 0,5% ;  rdg= 12,69V; n= 1 digit;

ls digit value= 0,01 V

Digital device with unknown precision

Precision is commonly considered 2 times the least significant digit

Example: digital counter measuring time, display format is xxxx.xx seconds. The least significant digit is 0,01 s => Δ= 0,02 s

uB = Δ / sqrt(12) = 0,02 / sqrt (12) = 0,0058 s = 5,8 ms

Time measurement by a stopwatch

A student will estimate his/her precision os measurement.

Example Δ= 0,3 s

uB = Δ / sqrt(12) = 0,3 / sqrt (12) = 0,087 s = 87 ms

Combined uncertainty for direct measurement – C type uncertainty

Is given by a formula

Combined uncertainty for a function